## Slide 1: A Randomized Comparison of the ACURATE neo versus the SAPIEN 3 Transcatheter Heart Valve Systems in Patients with Symptomatic Severe Aortic Stenosis
Jonas Lanz, Won-Keun Kim, Thomas Walther, Christof Burgdorf, Helge Möllmann, Axel Linke, Simon Redwood, Christian Thilo, Michael Hilker, Michael Joner, Holger Thiele, Lars Conzelmann, Lenard Conradi, Sebastian Kerber, Gerhard Schymik, Bernard Prendergast, Oliver Husser, Stefan Stortecky, Dik Heg, Peter Jüni, Stephan Windecker, Thomas Pilgrim
on behalf of the SCOPE I investigators

## Slide 2
I, Jonas Lanz, DO NOT have a financial interest/arrangement or affiliation with one or more organizations that could be perceived as a real or apparent conflict of interest in the context of the subject of this presentation.
Disclosure Statement of Financial Interest

## Slide 3: Background
TF TAVR has become an indispensable treatment option for patients with symptomatic severe aortic stenosis across all risk categories
The generalizability of outcomes observed in landmark trials comparing TAVR with SAVR to other commercial TAVR systems is limited by differences in device properties and the lack of head-to-head device comparisons
Iterations of the balloon-expandable SAPIEN THV system have been extensively investigated in several large-scale, high-quality RCTs and registries setting the current benchmark in terms of safety and efficacy
The ACURATE neo is a novel, self-expanding TAVR prosthesis associated with favorable outcomes in non-randomized studies

## Slide 4: Objective
To compare early safety and efficacy of the self-expanding ACURATE neo to the balloon-expandable SAPIEN 3 transcatheter heart valve system in patients with symptomatic severe aortic stenosis undergoing transfemoral TAVR

## Slide 5
Study Devices
 | ACURATE neoTM Aortic Valve System © 2019 Boston Scientific Corporation | SAPIEN 3TM Transcatheter Heart Valve System © 2019 Edwards Lifesciences Corporation
Frame | Nitinol | Cobalt-chromium
Leaflets | Porcine pericardium, supra-annular | Bovine pericardium, intra-annular
Expansion | Self-expanding (top-down) | Balloon-expandable
Recapturable | No | No
Valve sizes | S (23mm), M (25mm), L (27mm) | 23mm, 26mm and 29mm
Sheath inner diameter | 18-French | 14- and 16-French expandable
Paravalvular leakage reduction | Outer & inner skirt | Outer cuff & inner skirt
CE mark / FDA approval | Sep 2014 / No | Jan 2014 / Jun 2015
Notes: 4

## Slide 6
Study Design
SAPIEN 3
ACURATE neo
Heart team decision
TF TAVR
SAVR
Patients with severe aortic stenosis requiring intervention
Screening Log
Randomized controlled trial (730 patients)
1:1 Randomization
Follow-up:
at 30-days, 1 and 3 years
Primary endpoint:
Combined early safety & clinical efficacy at 30 days (VARC-2)
Notes: 5

## Slide 7: Eligibility Criteria
Age ≥ 75 years
Aortic valve area < 1 cm2
Clinical symptoms
(> NYHA I, angina, syncope)
Inoperable or at increased risk for SAVR based on risk scores and/or heart team recommendation
Aortic annulus dimensions and peripheral access suitable for either device
Major inclusion Criteria Major Exclusion Criteria
Congenital anomaly of aortic valve
Emergency procedures
LV-EF< 20%
Left-sided prosthetic valve
Concomitant planned procedure (except for PCI)
Stroke or myocardial infarction
(prior 30 days)
Planned non-cardiac surgery
(next 30 days)

## Slide 8: Primary Endpoint
Safety & clinical efficacy at 30 days based on VARC-2
All-cause mortality
All stroke (disabling and non-disabling)
Life-threatening or disabling bleeding
Major vascular complication
Coronary artery obstruction requiring intervention
Acute kidney injury (stage 2 or 3)
Re-hospitalization for valve-related symptoms or worsening CHF
Valve related dysfunction requiring repeat procedure
Valve-related dysfunction (echocardiography): mean Gradient ≥20 mmHg and EOA ≤ 0.9-1.1 cm2 and/or DVI < 0.35 AND/OR ≥ moderate regurgitation

## Slide 9: Statistical Methods
Stratified randomization (by STS-PROM category and site)
Non-inferiority design
Assumed primary endpoint event rate: 22%
Non-inferiority margin: 7.7% (risk-difference)
Power: 80%
Type I error (α) = 0.05 (one-sided)
365 patients per group
Primary analysis in intention-to-treat cohort, risk difference pooled over STS strata by Mantel-Haenszel (M-H) method

## Slide 10
Sponsor
  Clinical Department of Cardiology, University Hospital Bern, Switzerland
Data management & Monitoring
  University Hospital & Clinical Trials Unit, University of Bern, Switzerland
Statistics
  Clinical Trials Unit, University of Bern, Switzerland
Clinical Events Committee
  Cardiovascular European Research Center (CERC), Massy, France
Echocardiography Core Laboratory
  Medical Research Development, Hospital La Zarzuela, Madrid, Spain
Funder
  Boston Scientific, Marlborough, Massachusetts, USA
Trial Organization
Notes: 9

## Slide 11
Study Sites
20 European sites, 4 Nations: Switzerland (3), Germany (15), Netherlands (1), UK (1)
Study Site | Local principal investigator
Klinikum Augsburg | Christian Thilo, MD
Zentralklinik, Bad Berka | Stefan Richter, MD
Heart and Vascular Center, Bad Bevensen | Christof Burgdorf, MD
Kerckhoff Heart and Thorax Center, Bad Nauheim | Won-Keun Kim, MD Thomas Walther, MD
Cardio-vascular Center Bad Neustadt, | Sebastian Kerber, MD
St.-Johannes-Hospital, Dortmund | Helge Möllmann, MD
Heart Center, Dresden | Axel Linke, MD
Helios Klinik, Karlsruhe | Lars Conzelmann, MD
St. Vincentius-Kliniken, Karlsruhe | Alexander Würth, MD
Städtisches Klinikum, Karslruhe | Gerhard Schymik, MD
University Heart Center, Cologne | Stephan Baldus, MD
Heart Center, Leipzig | Holger Thiele, MD
German Heart Centre, Munich | Michael Joner, MD
University Medical Center, Regensburg | Michael Hilker, MD
University Medical Center, Utrecht | Pieter Stella, MD
St Thomas` Hospital, London | Simon Redwood, MD
Bern University Hospital, Bern | Thomas Pilgrim, MD
Lucerne Cantonal Hospital, Lucerne | Stefan Toggweiler, MD
University Hospital Zurich, Zurich | Maurizio Taramasso, MD
Notes: 10

## Slide 12: Patient Flow Chart
372 allocated to ACURATE neo
367 allocated to SAPIEN 3
Randomization
5 withdrawal of consent
0 lost-to-follow-up
369 TF TAVR initiated
363 received ACURATE neo
11 multiple valve implantation
2 conversion to SAVR
6 received SAPIEN 3
3 TF TAVR not initiated
(2 deaths, 1 infection)
367 Primary endpoint analyzed
358 alive
9 died
367 (99%) Clinical endpoints assessed
361 (97%) Echocardiography performed and analyzed
364 Primary endpoint analyzed
361 alive
3 died
364 (99%) Clinical endpoints assessed
363 (99%) Echocardiography performed and analyzed
3 withdrawal of consent
0 lost-to-follow-up
30-day Follow-up
Analysis
363 TF TAVR initiated
362 received SAPIEN 3
2 multiple valve implantation
1 received ACURATE neo
4 TF TAVR not initiated
(2 deaths, 1 withdrawal, 1 planned TA TAVR)
739 patients with severe, symptomatic aortic stenosis selected for TF TAVR by the Heart Team

## Slide 13: Baseline Characteristics (intention-to-treat)
367 Primary endpoint analyzed
358 alive
9 died
364 Primary endpoint analyzed
361 alive
3 died
Analysis
 | ACURATE neo (N = 372) | SAPIEN 3 (N = 367)
Demographics |  | 
Age - years (mean  SD) | 82.6  4.3 | 83.0  3.9
Female sex | 218 (59%) | 202 (55%)
Symptoms |  | 
NYHA classification III or IV | 287 (77%) | 268 (73%)
Risk assessment |  | 
STS-PROM score (median (interquartile range)) | 3.7 (2.6, 4.9) | 3.4 (2.6, 5.2)
STS-PROM score categories |  | 
low STS-PROM (<3%) | 134 (36%) | 136 (37%)
intermediate STS-PROM (≥3% and <8%) | 207 (55%) | 203 (55%)
high STS-PROM (≥8%) | 31 (8%) | 28 (8%)

## Slide 14: Baseline Imaging Characteristics (intention-to-treat)
367 Primary endpoint analyzed
358 alive
9 died
364 Primary endpoint analyzed
361 alive
3 died
 | ACURATE neo (N = 372) | SAPIEN 3 (N = 367)
Echocardiography |  | 
Aortic valve mean gradient - mmHg (mean  SD) | 42.9  17.2 | 41.5  15.1
Aortic valve area - cm2 (mean  SD) | 0.7  0.2 | 0.7  0.2
Left ventricular ejection fraction - % (mean  SD) | 56.4  11.1 | 57.1  10.7
Computed tomography |  | 
Aortic annulus perimeter - mm (mean  SD) | 75.7  5.2 | 75.9  5.1
Aortic annulus area - mm2 (mean  SD) | 439.1  59.6 | 442.9  60.3
Aortic valve calcification, moderate or severe | 286 (77%) | 286 (78%)
LVOT calcification, moderate or severe | 94 (25%) | 99 (27%)

## Slide 15: Procedural Characteristics
Procedure Time Contrast Volume

## Slide 16: Procedural Characteristics
367 Primary endpoint analyzed
358 alive
9 died
364 Primary endpoint analyzed
361 alive
3 died
Analysis
Pre-dilatation Post-dilatation

## Slide 17: Procedural Adverse Events (intention-to-treat)
367 Primary endpoint analyzed
358 alive
9 died
364 Primary endpoint analyzed
361 alive
3 died
Analysis
 | ACURATE neo (N = 369) | SAPIEN 3 (N = 363) | P value
Valve mal-positioning | 5 (1%) | 2 (1%) | 0.26
Implantation of multiple valves | 11 (3%) | 2 (1%) | 0.0119
Coronary artery obstruction requiring intervention | 0 (0%) | 0 (0%) | 1.00
Peri-procedural myocardial infarction | 1 (0.3%) | 1 (0.3%) | 0.96
Cardiac tamponade | 4 (1%) | 5 (1%) | 0.72
Annular rupture | 2 (1%) | 1 (0.3%) | 0.57
Left ventricular perforation | 1 (0.3%) | 0 (0%) | 0.32
Conversion to open heart surgery | 3 (1%) | 0 (0%) | 0.08
SAVR | 2 (1%) | 0 (0%) | 0.50
Immediate procedural death | 3 (1%) | 1 (0.3%) | 0.32
 | ACURATE neo (N = 369) | SAPIEN 3 (N = 363) | P value
Valve mal-positioning | 5 (1%) | 2 (1%) | 0.26
Implantation of multiple valves | 11 (3%) | 2 (1%) | 0.0119
Coronary artery obstruction requiring intervention | 0 (0%) | 0 (0%) | 1.00
Peri-procedural myocardial infarction | 1 (0.3%) | 1 (0.3%) | 0.96
Cardiac tamponade | 4 (1%) | 5 (1%) | 0.72
Annular rupture | 2 (1%) | 1 (0.3%) | 0.57
Left ventricular perforation | 1 (0.3%) | 0 (0%) | 0.32
Conversion to open heart surgery | 3 (1%) | 0 (0%) | 0.08
SAVR | 2 (1%) | 0 (0%) | 0.50
Immediate procedural death | 3 (1%) | 1 (0.3%) | 0.32

## Slide 18: Primary Endpoint Rates at 30 days
367 Primary endpoint analyzed
358 alive
9 died
364 Primary endpoint analyzed
361 alive
3 died
Analysis
ACURATE neo | SAPIEN 3
87/367 (23.7%) | 60/364 (16.5%)
Intention-to-treat

## Slide 19: Primary Endpoint
Primary analysis at 30 days (intention-to-treat)
Non-inferiority margin: 7.7%
-5% 0% 2.2% 7.1% 12.0% Risk difference (M-H)
ACURATE neo better SAPIEN 3 better
Upper limit of one-sided 95% CI: 12%
-5% 0% 2.2% 7.1% 12.0% Risk difference (M-H)
-5% 0% 2.2% 7.1% 12.0% Risk difference (M-H)
P value for non-inferiority: 0.42
-5% 0% 2.2% 7.1% 12.0% Risk difference (M-H)
ACURATE neo 23.7% | SAPIEN 3: 16.5%
-5% 0% 2.2% 7.1% 12.0% Risk difference (M-H)

## Slide 20: Primary Endpoint - Secondary Analyses at 30 days
 | ACURATE neo | SAPIEN 3 | Risk difference % |  | P value
 | No. of events/total no. (%) |  | (95%-CI) |  | 
Primary endpoint (superiority analysis) | 87/367 (23.7%) | 60/364 (16.5%) |  |  | 0.0156
Single components of primary endpoint |  |  |  |  | 
All-cause death | 9/367 (2.5%) | 3/364 (0.8%) |  |  | 0.09
Stroke (any) | 7/367 (1.9%) | 11/364 (3.0%) |  |  | 0.33
Life-threatening or disabling bleeding | 14/367 (3.8%) | 9/364 (2.5%) |  |  | 0.30
Major vascular complications | 29/367 (7.9%) | 20/364 (5.5%) |  |  | 0.21
Coronary artery obstruction requiring intervention | 0/367 (0%) | 0/364 (0%) |  |  | n/a
Acute kidney injury, stage 2 or 3 | 11/367 (3.0%) | 3/364 (0.8%) |  |  | 0.0340
Re-hospitalization for valve-related dysfunction or CHF | 4/367 (1.1%) | 5/364 (1.4%) |  |  | 0.72
Valve-related dysfunction requiring repeat procedure | 3/367 (0.8%) | 1/364 (0.3%) |  |  | 0.32
Valve-related dysfunction (echocardiography) | 35/361 (9.7%) | 17/363 (4.7%) |  |  | 0.0084
Chart: 
| 7.1
 | 0 | 0.5 | 1.5 | 2.5 | 3.5 | 4.5 | 5.5 | 6.5 | 7.5 | 8.5 | 10 | 10.5 | 11
 | 0 | 0.5 | 1.5 | 2.5 | 3.5 | 4.5 | 5.5 | 6.5 | 7.5 | 8.5 | 9.5 | 10.5 | 11.5
 | 0 | 0.5 | 1.5 | 2.5 | 3.5 | 4.5 | 5.5 | 6.5 | 7.5 | 8.5 | 10 | 10.5 | 11
 | ACURATE neo | SAPIEN 3 | Risk difference % |  | P value
 | No. of events/total no. (%) |  | (95%-CI) |  | 
Primary endpoint (superiority analysis) | 87/367 (23.7%) | 60/364 (16.5%) |  |  | 0.0156
Single components of primary endpoint |  |  |  |  | 
All-cause death | 9/367 (2.5%) | 3/364 (0.8%) |  |  | 0.09
Stroke (any) | 7/367 (1.9%) | 11/364 (3.0%) |  |  | 0.33
Life-threatening or disabling bleeding | 14/367 (3.8%) | 9/364 (2.5%) |  |  | 0.30
Major vascular complications | 29/367 (7.9%) | 20/364 (5.5%) |  |  | 0.21
Coronary artery obstruction requiring intervention | 0/367 (0%) | 0/364 (0%) |  |  | n/a
Acute kidney injury, stage 2 or 3 | 11/367 (3.0%) | 3/364 (0.8%) |  |  | 0.0340
Re-hospitalization for valve-related dysfunction or CHF | 4/367 (1.1%) | 5/364 (1.4%) |  |  | 0.72
Valve-related dysfunction requiring repeat procedure | 3/367 (0.8%) | 1/364 (0.3%) |  |  | 0.32
Valve-related dysfunction (echocardiography) | 35/361 (9.7%) | 17/363 (4.7%) |  |  | 0.0084
Chart: 
| 5 | 0.54 | -0.3 | 2.1 | 2.2999999999999998 | 1.3 | -1.1000000000000001 | 1.6 | 7.1
 | 0 | 0.5 | 1.5 | 2.5 | 3.5 | 4.5 | 5.5 | 6.5 | 7.5 | 8.5 | 9.5 | 10.5 | 11
 | 0 | 0.5 | 1.5 | 2.5 | 3.5 | 4.5 | 5.5 | 6.5 | 7.5 | 8.5 | 9.5 | 10.5 | 11.5
 | 0 | 0.5 | 1.5 | 2.5 | 3.5 | 4.5 | 5.5 | 6.5 | 7.5 | 8.5 | 9.5 | 10.5 | 11
 | ACURATE neo | SAPIEN 3 | Risk difference % |  | P value
 | No. of events/total no. (%) |  | (95%-CI) |  | 
Primary endpoint (superiority analysis) | 87/367 (23.7%) | 60/364 (16.5%) |  |  | 0.0156
Single components of primary endpoint |  |  |  |  | 
All-cause death | 9/367 (2.5%) | 3/364 (0.8%) |  |  | 0.09
Stroke (any) | 7/367 (1.9%) | 11/364 (3.0%) |  |  | 0.33
Life-threatening or disabling bleeding | 14/367 (3.8%) | 9/364 (2.5%) |  |  | 0.30
Major vascular complications | 29/367 (7.9%) | 20/364 (5.5%) |  |  | 0.21
Coronary artery obstruction requiring intervention | 0/367 (0%) | 0/364 (0%) |  |  | n/a
Acute kidney injury, stage 2 or 3 | 11/367 (3.0%) | 3/364 (0.8%) |  |  | 0.0340
Re-hospitalization for valve-related dysfunction or CHF | 4/367 (1.1%) | 5/364 (1.4%) |  |  | 0.72
Valve-related dysfunction requiring repeat procedure | 3/367 (0.8%) | 1/364 (0.3%) |  |  | 0.32
Valve-related dysfunction (echocardiography) | 35/361 (9.7%) | 17/363 (4.7%) |  |  | 0.0084
Chart: 
| 5 | 2.1 | 7.1
 | 0 | 0.5 | 1.5 | 2.5 | 3.5 | 4.5 | 5.5 | 6.5 | 7.5 | 8.5 | 10 | 10.5 | 11
 | 0 | 0.5 | 1.5 | 2.5 | 3.5 | 4.5 | 5.5 | 6.5 | 7.5 | 8.5 | 9.5 | 10.5 | 11.5
 | 0 | 0.5 | 1.5 | 2.5 | 3.5 | 4.5 | 5.5 | 6.5 | 7.5 | 8.5 | 10 | 10.5 | 11

## Slide 21: Primary Endpoint - Per Protocol Analyses
367 Primary endpoint analyzed
358 alive
9 died
364 Primary endpoint analyzed
361 alive
3 died
Analysis
Non-inferiority
Non-inferiority margin = 7.7%
-2.0% 0% 2.0% 7.0% 12.0% Risk difference (M-H)
Upper-limit one-sided 95%-CI:12.1%
P value for non-inferiority: 0.39
Superiority

## Slide 22: New Pacemaker Implantation
367 Primary endpoint analyzed
358 alive
9 died
364 Primary endpoint analyzed
361 alive
3 died
Analysis
Numbers refer to the cohort at risk
(patients with pacemaker at baseline excluded)

## Slide 23: Echocardiographic Valve Performance
Mean Gradient ≥20 mmHg AND
EOA ≤ 0.9-1.1 cm2 and/or DVI < 0.35
Paravalvular
Aortic Regurgitation
Notes: Echocardiographic performances;
This slide shows the two components that are part of echocardiographically determined valve-related dysfunction, which was part of the primary endpoint and the driving component for the differences in the primary endpoints rates beween the two devices.
Prosthetic aortic regurgitation as assessed by the echocardiography core laboratory occurred in 9.4% in the ACURATE neo and 2.8% in the SAPIEN 3 group.
Conversely Mean gradient above 20mmHg and reduced effective orifice area and/or doppler velocity index was more frequent with the SAPIEN 3 but not to an extent that outweights the difference in prosthetic aortic regurgitation.
ADJUST 0 0.001
At 30 days
mean Gradient ≥20 mmHg and EOA ≤ 0.9-1.1 cm2 and/or DVI < 0.35 OR
Check whether one patient each who has AR and valve stenosis
Add Proportion with Definiton of Gradient > 20 and EOA…
May consider better illustrating the echo endpoint
Reflect on reasoning of choosing these And OR connections

## Slide 24: Echocardiographic Valve Performance
Mean Gradient Effective Orifice Area

## Slide 25: Limitations
Not powered for individual clinical endpoints
Early primary endpoint limits evaluation of device differences in terms of long-term clinical outcomes
Single-blinded trial
visible differences in the stent frame preclude allocation concealment of echocardiography core laboratory
Lack of pre-assessment of aortic root CT angiographies by a central core laboratory

## Slide 26: Summary of Major Results
Non-inferiority of ACURATE neo versus SAPIEN 3 with respect to composite safety and efficacy endpoint at 30 days not met
Superiority of SAPIEN 3 with regard to composite safety and efficacy endpoint at 30 days in secondary analyses, driven by lower rates of paravalvular regurgitation and acute kidney injury (stage 2 or 3)
Higher rates of multiple valve implantation with ACURATE neo
Lower transvalvular gradients and larger effective orifice area with ACURATE neo
Low mortality, stroke and pacemaker rates with both devices

## Slide 27: Conclusions
ACURATE neo did not meet non-inferiority compared to the SAPIEN 3 device regarding the primary composite safety and efficacy endpoint at 30 days
Differences between the two TAVR devices were driven by moderate or severe paravalvular regurgitation and stage 2 or 3 acute kidney injury in favor of the SAPIEN 3 device
An early composite safety and efficacy endpoint proved useful in discriminating the performance of different TAVR systems