## Slide 1: Contemporary Outcomes with MitraClip™ (NTR/XTR) System in Primary Mitral Regurgitation: Results from the Global EXPAND Study
D. Scott Lim, MD
Wolfgang Rottbauer, Paul Mahoney, Ralph Stephan von Bardeleben, Matthew J. Price, Mathew Williams, Mustafa Ahmed, Bassem Chehab, Gilbert H. L. Tang, Annapoorna Kini, Jose L. Zamorano, Federico Asch, Francesco Maisano, and Saibal Kar - on behalf of the EXPAND Study Investigators
Professor of Medicine
University of Virginia, VA, USA

## Slide 2
D. Scott Lim, MD
Personal consulting:
Abbott, Edwards Lifesciences, Pipeline, Venus
Equity:
510Kardiac, Venus
My institution receives research grants on my behalf from:
Abbott, Boston Scientific, Corvia, Edwards Lifesciences, Keystone Heart, LivaNova, Medtronic, WL Gore
Disclosure Statement
EXPAND (NCT03502811) is funded by Abbott

## Slide 3
The MitraClip™ System was FDA approved in 2013 for the treatment of subjects with primary (or degenerative) mitral regurgitation who are at prohibitive risk for mitral valve surgery
Since the original EVEREST II trials, no new core-lab adjudicated data on subjects with primary MR have been reported
Background MitraClip™ for the Treatment of Primary MR
Notes: 2

## Slide 4
NTR
XTR
22mm
Improved delivery system
Longer arms, with an improved delivery catheter system
Background MitraClip™ Evolution - NTR/XTR
17mm
9mm
12mm
Notes: 3

## Slide 5: Primary Objective
To Report on:
Real World, Echo Core Lab and Clinical Events Committee Adjudicated Outcomes in Patients with Significant Primary Mitral Regurgitation Treated with the Next Generation MitraClip™ NTR and XTR Systems in the 1000+ Patient Global EXPAND Study
Notes: 4

## Slide 6: Secondary Objectives
Evaluate MR reduction outcomes as a function of mitral valve anatomic complexity
Confirm the safety (including single leaflet device attachments and leaflet injuries) and effectiveness of MR reduction associated with the MitraClip™ NTR and XTR systems

## Slide 7: Methods EXPAND Study Design
Design: Prospective, Multi-center, Single Arm, International, Post-Market, Real World, Observational Study Conducted in United States, Europe, and the Middle East
Intended Sample Size: A minimum of 1000 consecutive consented subjects receiving the MitraClip NTR/XTR at up to 60 sites
Follow-up: Baseline, Discharge, 30 days, 6 months, and 12 months
Key Outcome Measures: MR Severity, Major Adverse Events, Survival, Procedural Measures, Quality of Life Assessed using Kansas City Cardiomyopathy Questionnaire (KCCQ), and functional capacity assessed using New York Heart Association (NYHA) functional class
Data Adjudication: Clinical Events Committee (CEC) utilized for adjudication of adverse events; Echocardiographic Core Laboratory (ECL) utilized for adjudication of echocardiographic measures; An Independent Physician Committee utilized for evaluation of single leaflet device attachment (SLDA) and leaflet damage events
Notes: 6

## Slide 8: Methods EXPAND Study Design
1. Subjects with symptomatic MR (≥3+) as assessed by sites
2. Subjects provided consent for study participation
3. Subjects eligible to receive the MitraClip per the current approved indications for use in their respective geographies
1. Subjects participating in another clinical study that may impact the follow-up or results of this study.
INCLUSION CRITERIA
EXCLUSION CRITERIA
Notes: 7

## Slide 9
ECL MR Severity Assessment Methodology
MR Grade | Color Doppler Area Size | PV flow | Vena Contracta
1+, Mild | Central and small, (<4 cm2 or <10% of the LA area) | Systolic dominant | <0.5 cm
2+, Moderate | Central and moderate (4-6 cm2 or 10-30% of the LA area) | Diastolic dominant | 
3+, Moderate to Severe | Central, large (6-8 cm2 or 30-40% of the LA area) or eccentric reaching 1 PV. | All diastolic (systolic blunting) | ≥0.5 cm
4+, Severe | Eccentric, large (≥8 cm2 or ≥40% of the LA area), or eccentric reaching the second PV | Systolic flow reversal | 
A multiparametric algorithm adapted from the criteria recommended by the American Society of Echocardiography Guidelines employed for MR severity assessments consistent with the previous MitraClip trials (EVEREST II, REALISM).
MR Severity 3+ or 4+ (graded by 1 of 3 criteria) | 
Tier 1 | EROA ≥ 0.3 cm2 or PV systolic flow reversal
Tier 2 EROA 0.2 cm2 - <0.3 cm2 | With any 1 of the following: RV ≥ 45 ml/beat RF ≥ 40% VC width ≥ 0.5 cm
Tier 3 EROA not measured or <0.2 cm2 | With at least 2 of the following: RV ≥ 45 ml/beat RF ≥ 40% VC width ≥ 0.5 cm PISA radius > 0.9 cm, but CW of MR jet not done Large (≥ 6.0 cm) holosystolic jet wrapping around LA Peak E velocity ≥ 150 cm/s
Notes: 8

## Slide 10
ECL MV Complexity Assessment Criteria
Presence of at least one of the functional and anatomic features below at the baseline echocardiogram:
Primary jet outside of A2P2
Presence of more than one significant jet
Presence of an extremely wide jet
Small Valve
Calcification Landing Zone
Minimal leaflet tissue for attachment
Presence of severely degenerative leaflets or wide flail gaps or widths
Notes: 9
Federico provided some more clarity on the methodology used for assessing mixed etiology and the MV complexity – please see below. In nutshell his response states that the assessment methodology was consistent across the subjects, however didn’t involve any quantitative criteria (which he stated does not yet exists).
The MR etiology was determined based on which was considered the main etiology. If components of both functional and degenerative were present in a manner that we perceived neither was the dominant one, then we considered it Mixed. If we could not determine, then we called it non evaluable.
Presence of severely degenerative leaflet or wide flail gap or widths: Assessed visual impression of thickening leaflets with extensive redundancy
Small valve: <4cm2
Wide jet: A jet was considered wide when it was perceived that a single clip would not be able to correct the problem. This is a concept that comes from the early times of clipping, therefore, the definition is probably irrelevant (with the new clip types) at this point.

## Slide 11: Overall Study Subject Population
1041 Subjects Consented and Underwent MitraClip Implantation
422
Subjects with Primary or Mixed etiology
413
Subjects with Secondary MR Etiology
835 Subjects with Adequate Imaging for ECL Assessment
117 Subjects with Missing Baseline Images from Sites
89 Subjects not evaluable by ECL
Focus of the Present Analysis

## Slide 12: Echo Core Lab (ECL) Adjudication
Enrollment period: April 2018 through June 2019
2 ECLs utilized to ensure consistency of measurements across sites:
MR Etiology
MR Severity
MV Anatomic Characterization
LV Assessment
ECL -2
Utilized for detailed baseline mitral valve anatomic characterization and LV dimensional measurements over the study follow-up duration
ECL -1
Utilized for assessment of MR etiology (Baseline) and
MR severity (Baseline and Follow-Up)
Ensure consistent assessment with prior MitraClip trials (e.g. EVEREST II, REALISM, COAPT etc)
Notes: 11

## Slide 13: Baseline Characteristics
Demographics and Comorbidities | Value (N=422)
Age (years) | 79.5 ± 9.4
Male | 52.1% (220/422)
Body Mass Index (kg/m2) | 25.2 ± 4.8
STS Replacement Score* (%) | 7.3 ± 5.6
STS Repair Score (%) | 5.5 ± 5.4
Cardiac Arrhythmia | 60.5% (254/420)
Atrial Fibrillation | 92.5% (235/254)
Renal Failure | 28.2% (119/422)
Diabetes | 19.2% (80/416)
Dyslipidemia | 52.5% (219/417)
Hypertension | 79.5% (334/420)
Prior Heart Failure Hospitalization within 1 year | 43.2% (164/380)
Prior Myocardial Infarction | 13.5% (56/415)
* Continuous data presented as mean ± s.d.; Categorical data presented as a proportion of subjects where data was provided (missing data excluded)
** STS PROM formula has changed since EVEREST Trials. For reference, STS Replacement Mortality Risk (Mean±SD) for EVEREST II HRR was: 18.2± 8.0 (n=78); for EVEREST II REALISM was 11.1± 7.0 (n=628); and for Prohibitive Risk Primary MR Cohort was 13.2 ± 7.3% (n=127)

## Slide 14: ECL Adjudicated Baseline Echo Parameters for Subjects with Primary MR
ECL Measure | Value
Baseline MR of 3+ or 4+ | 66.4% (279/420)
Baseline MR of ≤2+* | 33.6% (141/420)
Effective Regurgitant Orifice Area (EROA, cm2) | 0.40 ± 0.21 (282)
Mean Mitral Inflow Velocity (cm/sec) | 47.0 ± 14.1 (342)
Baseline TR of 3+ or 4+ | 22.6% (84/372)
Left Ventricular Ejection Fraction (LVEF, %) | 62 ± 10 (390)
Left Ventricular End Systolic Volume (LVESV, mL) | 47.9 ± 30.4 (390)
Left Ventricular End Diastolic Volume (LVEDV, mL) | 121.0 ± 48.6 (390)
*These patients assessed as baseline MR Severity 3+/4+ by the sites.

## Slide 15: Procedural Outcomes vs. Previous Trials
 | PROHIBITIVE RISK PRIMARY MR 1,2 (N=127) | TVT Registry3 (N=2,952) | EXPAND PRIMARY MR (N=422)
Implant Rate % (n/N) [95% Confidence Interval] | 95.3%1 (121/127) | NA | 99.5% (420/422) [98.3%, 99.9%]
Acute Procedural Success (APS)* % (n/N) [95% Confidence Interval] | NA | 91.8% (2,709/2,952) Site-Reported | 94.5% (396/419) [91.9%, 96.5%] (ECL)
Fluoroscopy Time (min) Median [Inter-Quartile Range] | 39.02 [NA] | NA | 18.0 [11.3-27.0]
Procedure Time (min) Median [Inter-Quartile Range] | 1342 [NA] | NA | 82.0 [53.0- 120.0]
Length of Stay in Hospital for Index Procedure (days) Median [Inter-Quartile Range] | 2.9 ± 3.1 days1 (mean ± sd) | 2.0 [1.0–5.0] | 1.0 [1.0-3.0] (US)
*APS defined as successful implantation of the MitraClip® device with resulting MR severity of 2+ or less on discharge Echocardiogram (30-day echocardiogram was used if discharge is unavailable or uninterpretable). Subjects who die or undergo mitral valve surgery before discharge are considered to be an APS failure
1Lim et al. ACC 2018
2MitraClip Summary of Safety and Effectiveness for Prohibitive Risk DMR
3Sorajja et al. J Am Coll Cardiol 2017;70:2315–27)

## Slide 16: MitraClip Type Used in Subjects w/ Primary MR
Chart: 
| XTR Only | NTR Only | Both NTR & XTR
 | 0.4619 | 0.34762 | 0.19048
XTR was used more frequently to treat primary MR compared to NTR (p < 0.0001)

## Slide 17: Clip Usage by Baseline MR Severity
The baseline MR severity were statistically different across the 3 device groups
More XTR Clip usage (XTR-only or XTR and NTR) in subjects with baseline MR ≥3+
Baseline MR Severity | XTR Only (N=194) | NTR only (N=146) | XTR and NTR (N=80)
≤ 2+: Moderate1 | 26.3% (51/194) | 47.6% (69/145) | 25.3% (20/79)
≥ 3+: Severe1 | 73.7% (143/194) | 52.4% (76/145) | 74.7% (59/79)
1: p < 0.0001 Pearson’s Chi-square test
Notes: 16
More subjects with XTR (XTR only or XTR+NTR) have baseline MR 3+ or higher.

## Slide 18: Mitral Gradients (ECL Adjudicated) by MitraClip Size
 | XTR Only (N=194) | NTR only (N=146) | XTR and NTR (N=80) | Total (N= 420)
Pre Procedure Mitral Gradient, mmHg | 2.51 ± 1.35 (145) | 2.31 ± 1.22 (121) | 2.71 ± 1.62 (57) | 2.48 ± 1.40 (318)
30 day Mitral Gradient, mmHg | 3.51 ± 1.69 (156) | 3.89 ± 1.89 (121) | 3.99 ± 1.82 (66) | 3.74 ± 1.79 (337)
Pre-procedure gradient was not different across clip type groups (P =0.17)
MitraClip XTR was not associated with increased mitral valve gradient post procedure compared to MitraClip NTR.

## Slide 19: 30-Day Major Adverse Events*
All-cause Death | 2.4% (N=10)
Myocardial Infarction | 0.0% (N=0)
Peri-procedural MI | 0.0% (N=0)
Spontaneous MI | 0.0% (N=0)
Stroke | 1.2% (N=5)
Ischemic | 1.0% (N=4)
Hemorrhagic | 0.2% (N=1)
Non-elective CV surgery for device-related complications | 0.9% (N=4)
*CEC Adjudicated
Notes: Change to include 30 day mortality only.

## Slide 20: Device Related Leaflet Adverse Events
An independent multidisciplinary physician committee (including the ECL) reviewed and adjudicated all SLDA and leaflet injury events (including review of procedural and follow-up echoes as well as confirmation from surgery, etc.)
Overall, 8 subjects were confirmed to have an SLDA and 1 subject experienced a leaflet injury
*Adjudication for 3 SLDA events were inconclusive due to missing echo images at the time of the panel review, and not included in the final adjudication results.
 | Confirmed Events
Single Leaflet Device Attachment (SLDA) | 1.9% (N=8)*
Leaflet injury (leaflet tear or perforation) | 0.2% (N=1)
Notes: Below I summarized the details of the EXPAND Leaflet Events Committee which independently investigated the site reported mitral valve leaflet adverse events observed in EXPAND study.
One of the observations in EXPAND study is the inconsistencies in definitions used to identify and report mitral valve leaflet adverse events. Therefore, the goal of this committee was to adjudicate the events using all available imaging, clinical and surgical report, and to come up with standardized definitions for such leaflet-related adverse events and their echocardiographic characteristics.
The panel of expert independent physicians, comprised of interventional echocardiographers and interventional cardiologist (chaired by Dr. Federico Asch and includes Drs. Stephen H. Little, G. Burkhard Mackensen, Paul Grayburn, Paul Sorajja, Michael Rinaldi, Saibal Kar), reviewed the MitraClip NTR/XTR related suspected mitral valve leaflet injury (n=11) and single leaflet device attachment (SLDA) alone (n=24) cases reported by EXPAND Study sites. The adjudication results submitted/accepted in EuroPCR as LBT presentation (CONFIDENTIAL AND IN EMBGARGO) are as follows:
  4 out of 11 suspected leaflet injury cases were confirmed as leaflet tear.
  14 out of 24 suspected SLDA-alone cases were confirmed as SLDA.
The panel identified the specific echocardiographic characteristics of each observed leaflet events including leaflet perforation/tear, SLDA, chordal events (rapture and entrapment), partial leaflet gripping.
The standardized definitions for each mitral valve leaflet adverse events will be finalized and published in a scientific journal by the physician panel. The manuscript will include the EXPAND leaflet adverse event adjudication results, examples for each leaflet adverse event and their echocardiographic definitions for providing guidance to properly and consistently diagnose and report leaflet events with MitraClip.

## Slide 21: ECL Adjudicated MR Severity
Chart: 
| Baseline N=420 | Discharge N=403 | 30 Days N=365
MR 0+ | 0 | 0.377 | 0.277
MR 1+ | 0.055 | 0.491 | 0.592
MR 2+ | 0.281 | 0.102 | 0.104
MR 3+ | 0.331 | 0.022 | 0.022
MR 4+ | 0.333 | 0.008 | 0.005
MR ≤ 1+
86.9%
MR Reduction to None/Trace was achieved in 27.7% of primary MR subjects; MR ≤ 1+ was achieved in 86.9% and MR ≤ 2+ was achieved in 97.3% at 30 days follow up.
MR ≤ 2+
97.3%
MR = 0+
27.7%
Notes: OK to keep

## Slide 22: ECL Adjudicated MR Severity – by Baseline MR
Chart: 
| Baseline N=141 | Discharge N=133 | 30 days N=125 | Baseline N=279 | Discharge N=268 | 30 Days N=238
MR 0+ | 0 | 0.504 | 0.408 | 0 | 0.313 | 0.21
MR 1+ | 0.163 | 0.465 | 0.544 | 0 | 0.508 | 0.618
MR 2+ | 0.837 | 0.023 | 0.04 | 0 | 0.138 | 0.135
MR 3+ | 0 | 0.008 | 0.008 | 0.498 | 0.03 | 0.029
MR 4+ | 0 | 0 | 0 | 0.502 | 0.011 | 0.008
MR ≤ 2+
96.3%
MR ≤ 1+
95.2%
MR ≤ 2+
99.2%
MR ≤ 1+
82.8%
Baseline MR Severity ≤ 2+
(n=141)
Baseline MR Severity ≥ 3+
(n=279)
Notes: OK to keep

## Slide 23: ECL Adjudicated MR Severity
Chart: 
| Baseline N=279 | Discharge N=268 | Baseline N=123 | Discharge N=123
MR 0+ | 0 | 0.313 | 0 | 0.016
MR 1+ | 0 | 0.508 | 0 | 0.52
MR 2+ | 0 | 0.138 | 0.096 | 0.285
MR 3+ | 0.498 | 0.03 | 0.589 | 0.13
MR 4+ | 0.502 | 0.011 | 0.315 | 0.049
MR Reduction to ≤ 1+ at 30 days was numerically greater in EXPAND compared to Prohibitive Risk Cohort from prior EVEREST studies (Same ECL and methodology used)
EVEREST/REALISM Prohibitive Risk
Primary MR Cohort1 (n=123)
EXPAND Primary MR Subjects
w/ Baseline MR Severity ≥ 3+ (n=279)
MR ≤ 1+
82.1%
MR ≤ 2+
95.9%
52.0%
MR ≤ 1+
53.6%
MR ≤ 2+
82.1%
1Lim et al. ACC 2018
Notes: OK to keep

## Slide 24: Left Ventricle Dimensions and Volumes at 30 Days vs. Baseline
LV Remodeling
Early evidence of reduction in left ventricular dimensions and volumes; Larger, more clinically significant changes may be evident at longer term follow-up
Chart: LV Diameter (mm)
| Baseline 30 days LVESD | Baseline 30 days LVEDD
Baseline | 35.2 | 51.9
30 days | 33.2 | 49.3
P<0.0001
P<0.0001
Chart: LV Volume (mL)
| Baseline 30 days LVESV | Baseline 30 days LVEDV
Baseline | 47 | 120.6
30 days | 44.7 | 110
P<0.0001
P=0.01
Notes: Flail Gap (cm)
0.60 ± 0.32 (69)
0.45 ± 0.22 (36)
Prolapse Gap (cm)
0.41 ± 0.23 (69)
0.32 ± 0.18 (41)

## Slide 25: Quality of Life and Functional Capacity
Chart: 30 Day Change in KCCQ Score (N=354)
| Baseline | 30 Days
 | 50.17 | 70
* Pairwise comparison to Baseline
Chart: NYHA Class Change
| Baseline | Discharge | 30 Days | N=421 | N=373 | N=352
NYHA Class I | 0.048 | 0.153 | 0.315
NYHA Class II | 0.216 | 0.453 | 0.526
NYHA Class III | 0.641 | 0.37 | 0.15
NYHA Class IV | 0.095 | 0.024 | 0.009
P < 0.001
P < 0.001*

## Slide 26: Outcomes stratified by mitral valve anatomic complexity Echo Core-Lab Assessment

## Slide 27: ECL Adjudicated MV Complexity
Chart: Complexity of MV Anatomy (N=404)
| Non-Complex | Complex
 | 0.71 | 0.29
MV Complexity Criteria | % of Complex Anatomies (N=115)
Primary jet outside of A2P2 | 7.8% (9/115)
Presence of more than one significant jet | 0.9% (1/115)
Presence of an extremely wide jet | 29.6% (34/115)
Small Valve | 3.5% (4/115)
Calcification Landing Zone | 35.7% (41/115)
Minimal leaflet tissue for attachment | 5.2% (6/115)
Presence of severely degenerative leaflets or wide flail gaps or widths | 62.6% (72/115)

## Slide 28: Baseline MR Severity by MV Complexity
Subjects with complex anatomies had greater degrees of MR at baseline compared to subjects with non-complex anatomies
Baseline MR Severity | Subjects with Complex Anatomies (N=115) | Subjects with Non-Complex Anatomies (N=289)
2+ or less1 | 20.9% (24/115) | 37.7% (109/289)
3+ or higher1 | 79.1% (91/115) | 62.3% (180/289)
1: p = 0.0011

## Slide 29: MR Reduction – Complex vs. Non-Complex
Chart: 
| Baseline N=114 | 30 Days N=97 | Baseline N=289 | 30 Days N=253
MR 0+ | 0 | 0.196 | 0 | 0.277
MR 1+ | 0.017 | 0.598 | 0.065 | 0.612
MR 2+ | 0.191 | 0.175 | 0.31 | 0.083
MR 3+ | 0.261 | 0.031 | 0.357 | 0.02
MR 4+ | 0.531 | 0 | 0.268 | 0.008
Complex Anatomy (N=115)
Non-Complex Anatomy (N=289)
MR ≤ 1+
79.4%
MR ≤ 1+
88.9%
MR ≤ 2+
97.2%
MR ≤ 2+
96.9%

## Slide 30: NTR/XTR CLIP USE BASED ON mitral valve Anatomic measures Echo Core-Lab Assessment

## Slide 31: MV Anatomic Complexity
Type of clip used was not different across subjects with non-complex vs. complex MV anatomies
 | XTR Only (N=194) | NTR only (N=146) | XTR and NTR (N=80)
Non-Complex (n=289) | 71.8% (135/188) | 72.8% (99/136) | 68.7% (55/80)
Complex (n=115) | 28.2% (53/188) | 27.2% (37/136) | 31.3% (25/80)
Notes: 30

## Slide 32: XTR vs NTR Use
Prolapse Gap
Flail Gap
Chart: Prolapse Gap (mm)
| XTR | NTR
Series 1 | 4.24 | 3.27
Chart: Flail Gap (mm)
| XTR | NTR
Series 1 | 5.94 | 4.65
XTR was more often used in subjects with larger prolapse/flail gaps
MV Anatomic Characteristics
P<0.017
P<0.024
Notes: Flail Gap (cm)
0.60 ± 0.32 (69)
0.45 ± 0.22 (36)
Prolapse Gap (cm)
0.41 ± 0.23 (69)
0.32 ± 0.18 (41)

## Slide 33
Chart: Subjects with MR Reduction ≥ 2 Grades (%)
| XTR-only (n=45) | NTR-only (n=30)
MR reduction of 2 grade | 86.7 | 66.7
Greater MR reduction achieved using the XTR clip in Complex Anatomies (p = 0.0383)
30 day MR Reduction of ≥ 2 Grades
P=0.0383
Complex Anatomy
Chart: Subjects with MR Reduction ≥ 2 Grades (%)
| XTR-only (n=117) | NTR-only (n=88)
MR reduction of 2 grade | 67.5 | 64.8
Non-Complex Anatomy
P=0.6802
Notes: Flail Gap (cm)
0.60 ± 0.32 (69)
0.45 ± 0.22 (36)
Prolapse Gap (cm)
0.41 ± 0.23 (69)
0.32 ± 0.18 (41)

## Slide 34: Limitations
Real world nature of the study results in variability in the quality of echocardiographic acquisitions and hence certain measurements were not available for all patients
Enrollment was based on site interpretation of MR Severity rather than prospective core lab adjudication
Only acute procedure outcomes presented; Longer term follow-up ongoing

## Slide 35: Conclusions
This study represents the first contemporary report of ECL and CEC adjudicated 30 day clinical outcomes in patients with primary MR treated with the next generation NTR and XTR MitraClips
Approximately one-third of patients had a complex mitral valve anatomy that reflects the difference in patients treated in the real world in comparison to past clinical trials.
Results show that MR ≤ 1+ is being achieved more often with MitraClip™ NTR and XTR than previously observed in EVEREST II trials
MitraClip™ XTR was associated with greater MR reduction compared to NTR, in more complex anatomies
Notes: Can mention this one too:
Differences in methodology used for assessing MR Severity by the echo core lab and the sites (also US vs. OUS) has been reported.

## Slide 36: Summary
The EXPAND study confirms the safety and effectiveness of the next generation MitraClip NTR/XTR system in Primary MR patients in a contemporary, real-world setting

## Slide 37: Thank you

## Slide 38: APPENDIX

## Slide 39: Baseline Characteristics (ii)
* Continuous data presented as mean ± s.d.; Categorical data presented as a proportion of subjects where data was provided (missing data excluded)
** One subject with transcatheter repair also underwent a surgical procedure, and counted as one under surgery.
***One subject had both prior aortic and mitral valve procedures, and counted as one under mitral repair.
Prior Interventions | Value (N=422)
Permanent Pacemaker | 12.1% (51/422)
Cardiac Resynchronization Therapy (CRT or CRT-D) | 3.6% (15/422)
Intracardiac Defibrillation (ICD or CRT-D) | 5.9% (25/422)
Prior Valve Procedure (Surgery or Transcatheter) | 15.4% (65/422)
Prior Mitral Valve Procedure (Surgery or Transcatheter) | 6.6% (28/422)
Prior Surgical Mitral Valve Repair or Annuloplasty | 3.3% (14/422)
Prior Transcatheter Mitral Valve Repair** | 3.3% (14/422)
Prior Aortic Valve Procedure (Surgery or Transcatheter)*** | 8.8% (37/422)
Prior Coronary Revascularization | 30.8% (130/422)
Prior CABG | 11.4% (48/422)
Prior PCI | 24.7% (104/422)

## Slide 40: ECL Adjudicated MR Severity – US vs. OUS
No significant difference in MR reduction outcomes between US and OUS Subjects
Chart: 
| Baseline N=209 | Discharge N=203 | 30 Days N=184 | Baseline N=213 | Discharge N=200 | 30 Days N=181
MR 0+ | 0 | 0.384 | 0.245 | 0 | 0.37 | 0.309
MR 1+ | 0.024 | 0.478 | 0.587 | 0.085 | 0.505 | 0.597
MR 2+ | 0.268 | 0.113 | 0.13 | 0.29 | 0.09 | 0.077
MR 3+ | 0.301 | 0.02 | 0.033 | 0.357 | 0.025 | 0.011
MR 4+ | 0.407 | 0.005 | 0.005 | 0.258 | 0.01 | 0.006
MR ≤ 1+
83.2%
MR ≤ 2+
98.3%
US Subjects (N=208)
OUS Subjects (N=207)
MR ≤ 2+
96.2%
MR ≤ 1+
90.6%
Notes: OK to keep

## Slide 41: ECL Adjudicated Baseline Echo Parameters – Mean ± SD, Range
Effective Regurgitant Orifice Area (EROA, cm2) Mean ± SD (n) Range (Min, Max) | 0.40 ± 0.21 (282) (0.06, 1.52)
Mean Mitral Gradient (mean MVG, mmHg) Mean ± SD (n) Range (Min, Max) | 2.47 ± 1.36 (323) (0.38, 10.89)
Left Ventricular Ejection Fraction (LVEF, %) Mean ± SD (n) Range (Min, Max) | 62 ± 10 (390) (15.44, 81.63)
Left Ventricle End Systolic Diameter (LVESD, mm) Mean ± SD (n) Range (Min, Max) | 35.4 ± 7.7 (401) (20.9, 62.5)
Left Ventricle End Diastolic Diameter LVEDD, mm) Mean ± SD (n) Range (Min, Max) | 52.0 ± 6.6 (403) (35.6, 79.2)
Left Ventricle End Systolic Volume (LVESV, mL) Mean ± SD (n) Range (Min, Max) | 47.9 ± 30.4 (390) (11.89, 278.49)
Left Ventricle End Diastolic Volume (LVEDV, mL) Mean ± SD (n) Range (Min, Max) | 121.0 ± 48.6 (390) (41.25, 376.61)

## Slide 42: Length of Stay – US vs. OUS
 | US Subjects (N=209) | OUS Subjects (N=213)
Length of Stay in Hospital for Index Procedure (days) Mean ± SD (n) Median [Inter-Quartile Range] | 2.7 ± 3.5 (209) 1.0 (1.0, 3.0) | 8.0 ± 6.9 (213) 6.0 (4.0, 9.0)
Longer length of stay is reported for OUS subjects compared to US (p<0.0001)

## Slide 43: 30 Day MAE* vs. Previous Trials
 | Prohibitive Risk Primary MR Cohort1 (N=127) | EVEREST II HRR2 (N=78) | EVEREST II REALISM3 (N=628) | EXPAND Primary MR (N=420)
All-cause Death | 6.3% (8) | 7.7% (6) | 4.1% (26) | 2.4% (10)
Myocardial Infarction | 0.8% (1) | 2.6% (2) | 0.5% (3) | 0.0% (0)
Peri-procedural MI | NA | 0.0% (0) | 0.0% (0) | 0.0% (0)
Spontaneous MI | NA | 1.3% (1) | 0.5% (3) | 0.0% (0)
Stroke | 2.4% (3) | 2.6% (2) | 2.1% (13) | 1.2% (5)
Ischemic | NA | 2.6% (2) | NA | 1.0% (4)
Hemorrhagic | NA | 0.0% (0) | NA | 0.2% (1)
Non-elective CV surgery for device-related complications | 0.8% (1) | 0.0% (0) | 1.3% (8) | 0.9% (4)
1Lim et al. Prohibitive Risk Cohort. ACC 2018,
2Whitlow et al. EVEREST II HRS. JACC 2012,
3Feldman et al. REALISM HR data. EuroPCR 2015
*CEC Adjudicated
PROHIBITIVE RISK COHORT DATA ADDED

## Slide 44: MV Anatomic Complexity / TTE
For primary MR subjects with complex MV anatomy:
  Calcification Landing Zone and Presence of severely degenerative leaflets or wide flail gap are statistically different across the three groups of device use
  More subjects with NTR have Calcification Landing Zone
  More subjects with XTR present severely degenerative leaflets or wide flail gaps
 | XTR Only (N=194) | NTR only (N=146) | XTR and NTR (N=80) | P-value
Primary jet outside of A2P2 | 4/53 (7.5%) | 1/37 (2.7%) | 4/25 (16.0%) | 0.1601
Presence of more than one significant jet | 1/53 (1.9%) | 0/37 (0.0%) | 0/25 (0.0%) | 1.0002
Presence of a wide jet | 16/53 (30.2%) | 8/37 (21.6%) | 10/25 (40.0%) | 0.2961
Small Valve | 1/53 (1.9%) | 2/37 (5.4%) | 1/25 (4.0%) | 0.6832
Calcification Landing Zone | 11/53 (20.8%) | 20/37 (54.1%) | 10/25 (40.0%) | 0.0051
Minimal leaflet tissue for attachment | 5/53 (9.4%) | 1/37 (2.7%) | 0/25 (0.0%) | 0.2602
Presence of severely degenerative leaflets or wide flail gaps or widths | 39/53 (73.6%) | 17/37 (45.9%) | 16/25 (64.0%) | 0.0281
  1: Pearson’s Chi-square test
  2: Fisher’s exact test
Notes: 43
Fix the percentages

## Slide 45
Chart: 
| Baseline N=53 | 30 Days N=45 | Baseline N=37 | 30 Days N=30 | Baseline N=135 | 30 Days N=118 | Baseline N=98 | 30 Days N=89
MR 0+ | 0 | 0.267 | 0 | 0.233 | 0 | 0.237 | 0 | 0.304
MR 1+ | 0 | 0.578 | 0.054 | 0.5 | 0.067 | 0.61 | 0.081 | 0.629
MR 2+ | 0.132 | 0.111 | 0.351 | 0.233 | 0.237 | 0.111 | 0.414 | 0.056
MR 3+ | 0.321 | 0.044 | 0.162 | 0.034 | 0.363 | 0.025 | 0.343 | 0.011
MR 4+ | 0.547 | 0 | 0.433 | 0 | 0.333 | 0.017 | 0.162 | 0
30 Day MR Severity XTR vs. NTR
Complex Anatomy
Non-Complex Anatomy
XTR-only
NTR-only
XTR-only
NTR-only
MR ≤ 2+
95.6%
MR ≤ 1+
84.5%
MR ≤ 2+
95.8%
MR ≤ 1+
84.7%
MR ≤ 2+
98.9%
MR ≤ 1+
93.3%
MR ≤ 2+
96.6%
MR ≤ 1+
73.3%
XTR clip resulted more often in MR reduction ≤1+ in complex anatomies, whereas NTR clip achieved more often MR reduction to ≤1+ in non-complex anatomies, although the results are not statistically significant

## Slide 46: Procedural Outcomes vs. Previous Trials
 | EXPAND PRIMARY MR | EVEREST II HRR1 | EVEREST II REALISM2 | TVT Registry3
Implant Rate % (n/N) [95% Confidence Interval] | 99.5% (420/422) [98.3%, 99.9%] | 96.2% (75/78) | 96.0% (603/628) | NA
Acute Procedural Success (APS)* % (n/N) [95% Confidence Interval] | 94.5% (396/419) [91.9%, 96.5%] (ECL) | 71.8% (56/78) (ECL) | 84.8% (518/611) (ECL) | 91.8% (2,709/2,952) Site-Reported
Device Time (min) Median [Inter-Quartile Range] | 48.0 [33.0-76.0] | NA | 94.0 [NA] | NA
Fluoroscopy Time (min) Median [Inter-Quartile Range] | 18.0 [11.3-27.0] | NA | 33.0 [NA] | NA
Procedure Time (min) Median [Inter-Quartile Range] | 82.0 [53.0- 120.0] | NA | 126.0 [NA] | NA
Length of Stay in Hospital for Index Procedure (days) Median [Inter-Quartile Range] | 1.0 [1.0-3.0] (US) | 2.0 [NA] | 2.0 [NA] | 2.0 [1.0–5.0]
*APS defined as successful implantation of the MitraClip® device with
resulting MR severity of 2+ or less on discharge Echocardiogram (30-day
echocardiogram is used if discharge is unavailable or uninterpretable).
Subjects who die or undergo mitral valve surgery before discharge are
considered to be an APS failure
1Whitlow et al. EVEREST II HRS. JACC 2012
2Feldman et al. REALISM HR data. EuroPCR 2015
2Sorajja et al. J Am Coll Cardiol 2017;70:2315–27)

## Slide 47: ECL Adjudicated MR Severity
Chart: 
| Baseline N=279 | 30 Days N=238 | Baseline N=78 | 30 days N=70 | Baseline N=587 | 30 days N=587
MR 0+ | 0 | 0.21 | 0 | 0 | 0 | 1.2
MR 1+ | 0 | 0.618 | 0 | 48.6 | 1.4 | 53.7
MR 2+ | 0 | 0.135 | 1.3 | 24.3 | 18.4 | 35.1
MR 3+ | 0.498 | 0.029 | 76.9 | 22.9 | 51.6 | 8.5
MR 4+ | 0.502 | 0.008 | 21.8 | 4.3 | 28.6 | 1.5
MR Reduction to ≤ 1+ at 30 days was numerically greater in EXPAND compared to prior studies (Same ECL used in all 3 studies for MR Severity assessment)
EVEREST II
HRR (n=78)
EVEREST II
REALISM HR (n=587)
EXPAND Primary MR Subjects
w/ Baseline MR Severity ≥ 3+ (n=279)
MR ≤ 1+
82.8%
MR ≤ 2+
96.3%
MR ≤ 1+
48.6
MR ≤ 1+
54.9
MR ≤ 2+
72.9
48.6
MR ≤ 2+
88.8
1.2
Notes: OK to keep

## Slide 48: Mitral Valve Annular Dimensions at 30 Days vs. Baseline
Significant reduction in annular dimensions within 30 days of treatment
Chart: Annular Dimension (mm)
| Baseline 30 days | Baseline 30 days | Baseline 30 days | Baseline 30 days
Baseline | 28.8 | 32.2 | 28.9 | 31.6
30 days | 25.9 | 29 | 25.9 | 28.8
P<0.0001
P<0.0001
P<0.0001
P<0.0001
AP1 Systolic
Annular D.
AP1 Diastolic
Annular D.
CC2 Systolic Annular D.
CC2 Diastolic Annular D.
1AP: Anterior Posterior
2CC: Commissure to Commissure
Mitral Valve Remodeling of the Annulus
Notes: Flail Gap (cm)
0.60 ± 0.32 (69)
0.45 ± 0.22 (36)
Prolapse Gap (cm)
0.41 ± 0.23 (69)
0.32 ± 0.18 (41)

## Slide 49: MitraClip Use in Subjects w/ Primary MR
Chart: Clip Size Use by # of Clips
| 1 Clip | 2 Clips | 3 or more clips | 52% | 43% | 5% | (217 pts) | (180 pts) | (23 pts)
XTR Only | 0.28571 | 0.15981 | 0.016667
NTR Only | 0.23095 | 0.10714 | 0.0095238
Both NTR & XTR | 0 | 0.1619 | 0.028571
Chart: Clip Size Use in Primary MR
| XTR Only | NTR Only | Both NTR & XTR
 | 0.4619 | 0.34762 | 0.19048
XTR was used more frequently to treat primary MR compared to NTR (p < 0.0001)
37.8%
24.9%
37.3%
30.4%
17.9%
51.8%

## Slide 50: Mitral Gradients (ECL Adjudicated) by MitraClip Size
 | XTR Only (N=194) | NTR only (N=146) | XTR and NTR (N=80) | Total (N= 420)
Pre Procedure Mitral Gradient, mmHg | 2.51 ± 1.35 (145) | 2.31 ± 1.22 (121) | 2.71 ± 1.62 (57) | 2.48 ± 1.40 (318)
Discharge Mitral Gradient, mmHg | 3.48 ± 1.92 (173) | 3.96 ± 1.81 (124) | 4.02 ± 1.68 (73) | 3.75 ± 1.87 (363)
30 day Mitral Gradient, mmHg | 3.51 ± 1.69 (156) | 3.89 ± 1.89 (121) | 3.99 ± 1.82 (66) | 3.74 ± 1.79 (337)
1 Clip Only | 61.9% (120/194) | 66.4% (97/146) | 0.0% (0/80) | 51.7% (217/420)
2 Clips | 34.5% (67/194) | 30.8% (45/146) | 85% (68/80) | 42.9% (180/420)
3+ Clips | 3.6% (7/194) | 2.7% (4/146) | 15% (12/80) | 5.0% (21/420)
Pre-procedure gradient was not different across clip type groups (P =0.17)
MitraClip XTR was not associated with increased mitral valve gradient post procedure compared to MitraClip NTR.