## Slide 1: Primary Results From the Evolut Low Risk Bicuspid Study
Basel Ramlawi, MD, FACC
Valley Health System, Winchester, Virginia For the Evolut Low Risk Bicuspid Investigators

## Slide 2: Disclosure Statement of Financial Interests
Within the past 12 months, I have had a financial interest/arrangement or affiliation with the organization(s) listed below.
Financial Relationship | Company
Research grants, consulting fees | AtriCure, Liva Nova, Medtronic
Medtronic personnel performed all statistical analyses and assisted with the graphical display of the data presented. | 

## Slide 3: Background
TAVR with Evolut supra-annular self-expanding valves has demonstrated excellent outcomes in tricuspid aortic stenosis and is currently approved in the US for patients across all risk classes.
Patients with bicuspid aortic stenosis have been generally excluded from prior TAVR trials, due to concerns of:
  Asymmetric calcification, elliptical shape, potential incomplete valve expansion, procedural technical concerns, etc.
  Annular vs. supra-annular measurements
There have been no prospective studies assessing TAVR in low risk patients with bicuspid aortic stenosis.

## Slide 4: Objective
To assess the safety and efficacy of TAVR in patients with bicuspid aortic valve stenosis and low surgical risk

## Slide 5
The Mount Sinai Medical Center New York, NY
St. Francis Hospital, Roslyn, NY
Abbott Northwestern Minneapolis, MN
Beth Israel Deaconess Mass General Boston, MA
Northwell Health, Manhasset, NY
Mercy Medical Center– Iowa Heart Des Moines, IA
Morton Plant Clearwater, FL
Baylor Heart & Vascular Dallas, TX
Houston Methodist DeBakey Heart & Vascular Center Houston, TX
St. Lukes/Aurora Milwaukee, WI
Piedmont Heart Institute Atlanta, GA
Case Medical Center Cleveland, OH
OhioHealth/Riverside Methodist Columbus, OH
Valley Health Winchester, VA
University of Michigan Ann Arbor, MI
Yale New Haven Hospital, New Haven, CT
Oregon Health & Science University Portland, OR
Tallahassee Research Institute Tallahassee, FL
Abrazo Arizona Heart Hospital Phoenix, AZ
Lee Memorial Health System Ft. Myers, FL
Los Robles Hospital Thousand Oaks, CA
Spectrum Health Grand Rapids, MI
Participating Sites
University Pittsburg Pinnacle Health Harrisburg, PA
Scripps Memorial La Jolla, CA

## Slide 6: Study Administration
Principal Investigators: John Forrest, Basel Ramlawi
Executive Committee: John Forrest, Jeffrey Popma, Basel Ramlawi, Michael Reardon
Screening Committee: G. Michael Deeb (Chair), John Forrest, Jeffrey Popma, Basel Ramlawi, Michael Reardon, Steven Yakubov
Echo Core Laboratory: Jae Oh, Mayo Clinic, Rochester, MN
Data & Safety Monitoring Board: David Faxon (Chair), William Holman, John Lopez, Scott Kasner, John Orav
Clinical Events Committee: Cliff Berger, Scott Bortman, Manish Chauhan, Donald Cutlip, Torin Fitton, Eli Gelfand, David Grossman, Claudia Hochberg (Chair), Carey Kimmelstiel, Daniel Kramer, Megan Leary, Robert Rodriguez, Sanjay Samy, Jonathan Silver, Gregory Smaroff, David Thaler, Jonathan Waks, Sergio Waxman, David Weiss, Jeffrey Veluz
Sponsor: Medtronic, Minneapolis, MN
5

## Slide 7: Study Methods
Multicenter, prospective, interventional, single-arm study
Baseline MSCT to confirm bicuspid morphology
Patient eligibility reviewed by local Heart Team & Screening Committee
Implant Procedure
  Annular sizing recommended for all patients
  Pre-TAVR balloon valvuloplasty strongly encouraged
CEC adjudicated all endpoint-related adverse events
Hemodynamics centrally assessed by echocardiographic core laboratory
Patient follow-up planned for 10 years
6

## Slide 8: Valves Studied
Evolut R Valve [43%] | Evolut PRO Valve [57%]
23, 26, 29 and 34 mm | 23, 26 and 29 mm

## Slide 9: Key Inclusion Criteria
Bicuspid aortic valve anatomy confirmed by MSCT
Symptomatic and asymptomatic severe AS1
A predicted risk of 30-day mortality <3% per multidisciplinary local Heart Team assessment
1Nishimura RA, et al. Circulation. 2014;129:2440-92.

## Slide 10: Key Exclusion Criteria
Age < 60 years
Multivessel coronary artery disease (SYNTAX score >22)
Ascending aorta diameter > 4.5 cm
Aortopathy requiring surgical intervention
Prohibitive LVOT calcification
Trileaflet aortic valve on MSCT
Anatomic dimensions outside recommended range
  SOV (≥ 25 mm)
  Annulus (18 to 30 mm)

## Slide 11: Study Endpoints
Primary Safety Endpoint
All-cause mortality or disabling stroke at 30 days
Primary Efficacy Endpoint
Device Success
  Absence of procedural mortality AND
  Correct position of 1 valve in the proper anatomical location AND
  Absence of > mild aortic valve regurgitation
10

## Slide 12: Patient Flow
11
Patients screened (n=222)
Excluded (n=72):
Exclusion criteria met (n=60)
Patient withdrawal (n=2)
Physician withdrawal (n=3)
Other (n=7)
Attempted procedure (n=150)
Aborted procedure (n=1)
Implanted TAV (n=149)
30-Day follow-up (n=147)
Death (n=1)
Discharged with TAV (n=148)
Conversion to surgery (n=1)

## Slide 13: Reasons for Study Exclusion
Reason | N = 60
Anatomical reasons | 46
Tricuspid aortic valve | 17
Aortic root dimensions: annular perimeter/diameter | 15
Mean ascending aorta > 45 mm | 9
Aortic root dimensions: SOV diameter | 4
Prohibitive LVOT calcification | 1
Risk of mortality outside protocol (> low risk) | 5
Contraindication for placement of bioprosthetic valve | 3
Allergies | 2
Did not meet severe AS criteria | 2
Age less than 60 years | 1
Other condition excluding from study per investigator | 1

## Slide 14: RESULTS

## Slide 15: Baseline Clinical Characteristics
Mean ± SD or no. (%) | N = 150
Age, years | 70.3 ± 5.5
Male sex | 78 (52.0)
Body surface area, m2 | 1.9 ± 0.2
STS PROM, % | 1.4 ± 0.6
NYHA Class III or IV | 41 (27.3)
Peripheral arterial disease | 14 (9.3)
Chronic lung disease/ COPD | 26 (17.7)
Prior coronary artery bypass grafting | 2 (1.3)
Mean gradient, mm Hg | 48.0 ± 16.1
Aortic valve area, cm2 | 0.8 ± 0.2

## Slide 16: Bicuspid Valve Sievers Subtypes
Type 0 N = 14 | Type I Left-Right N = 107 | Type I Right-Non N=27 | Type I Non-Left N= 2
No patients had Sievers Type 2.
9.3%
71.3%
18.0%
1.3%

## Slide 17: Procedural Characteristics
No. (%) | N = 150
General anesthesia | 95 (63.3)
Iliofemoral access | 147 (98.7)
Embolic protection | 45 (30.0)
Pre-TAVR balloon dilation | 137 (91.3)
Post-TAVR balloon dilation | 55 (36.9)
> 1 valve implanted | 5 (3.3)
No. (%) | N = 150
Resheath or recapture | 49 (32.9)
Implanted valve size | 
23 mm | 0 (0.0)
26 mm | 32 (21.5)
29 mm | 55 (36.9)
34 mm* | 62 (41.6)
*Only Evolut R

## Slide 18: Primary Endpoint
Chart: 
| 0 | 0.1 | 1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30
TAVR | 0 | 0 | 0 | 0.006667 | 0.006667 | 0.006667 | 0.006667 | 0.006667 | 0.006667 | 0.006667 | 0.006667 | 0.006667 | 0.006667 | 0.013333 | 0.013333 | 0.013333 | 0.013333 | 0.013333 | 0.013333 | 0.013333 | 0.013333 | 0.013333 | 0.013333 | 0.013333 | 0.013333 | 0.013333 | 0.013333 | 0.013333 | 0.013333 | 0.013333 | 0.013333 | 0.013333 | 0.013333 | 0.013333
1.3%
Death or Disabling Stroke
No. at Risk | 150 | 143
Days Post Procedure

## Slide 19: Outcomes at 30 Days
No. of patients (KM estimates as %) | N = 150
All-cause mortality or disabling stroke | 2 (1.3)
All-cause mortality | 1 (0.7)
Disabling stroke | 1 (0.7)
Non-disabling stroke | 5 (3.3)
Major vascular complication | 2 (1.3)
Aortic dissection | 0 (0.0)
Annular rupture | 0 (0.0)
Permanent pacemaker* | 22 (14.7)
Permanent pacemaker† | 22 (15.1)
Coronary artery obstruction | 1 (0.7)
*Includes patients with baseline permanent pacemaker. †Excludes patients with baseline permanent pacemaker.

## Slide 20: Additional Outcomes at 30 Days
No. of patients (%) | 
Device success | 141/148 (95.3)
Patient-prosthesis mismatch* | 
None | 115 (87.1)
Moderate | 10 (7.6)
Severe | 7 (5.3)
Mean gradient > 20 mm Hg | 2/146 (1.4)
*N=132 patients per VARC-2.

## Slide 21
Chart: 
| Baseline | Post Procedure | 30 Days
EOA | 0.8 | 2.2 | 2.3
Mean Gradient | 48 | 9 | 7.6
Valve Hemodynamics
EOA | 135 | 124 | 132
AVG | 148 | 146 | 146
No. of Echos

## Slide 22
Chart: 
| Baseline | Post Procedure | 30 Days
Type 0 EOA | 0.7 | 2.3 | 2.1
Type 1 EOA | 0.8 | 2.2 | 2.3
Type 0 AVG | 48.4 | 8.8 | 7.2
Type 1 AVG | 48 | 9.1 | 7.6
Valve Hemodynamics by Subtypes
Type 0
Type 1
21

## Slide 23
Patients w/Echo (%)
Total Aortic Valve Regurgitation at 30 Days
Implant population. Core lab assessments.
Chart: 
| All Patients N = 146 | Type 0 N = 13 | Type I N = 133
None/Trace | 0.596 | 0.846 | 0.571
Mild | 0.404 | 0.154 | 0.429
Moderate | 0 | 0 | 0
Severe | 0 | 0 | 0

## Slide 24
Chart: 
| N=150 | N=146 | Baseline | 30 Days
NYHA I | 0.02 | 0.753
NYHA II | 0.707 | 0.219
NYHA III | 0.267 | 0.027
NYHA IV | 0.007 | 0
Proportion of Patients (%)
New York Heart Association

## Slide 25: Discussion
Study performed at high volume expert centers
Non-randomized study design
Standardized implant technique
  Annular sizing
  Pre-TAVR balloon dilation
Rigorous adherence to patient selection parameters

## Slide 26: Summary
TAVR with Evolut supra-annular self-expanding valve in low-risk bicuspid patients achieved excellent early results:
  Annular sizing achieved 95.3% device success
  Low mortality and stroke at 30 days (1.3%)
  Low rates of PVL (no moderate/severe)
  Consistent hemodynamics across Sievers Classification
Patients will be followed for 10 years

## Slide 27: Clinical Implications
In low-risk AS patients with bicuspid morphology, TAVR with the Evolut platform is a safe and effective treatment option.
The choice between TAVR and surgery should be based on a multidisciplinary heart team discussion that includes anatomic, clinical and patient social factors.